## Compliance rate
Table 5：COVID-19 Positive Reference Product Compliance Test Results
Batch.NO. | 5 measurement results （IgM/IgG）
20170104 | +/- | +/+ | +/- | -/+ | -/+
20170405 | 5 measurement results （IgM/IgG）
| +/- | +/+ | +/- | -/+ | -/+
20170502 | 5 measurement results （IgM/IgG）
| +/- | +/+ | +/- | -/+ | -/+
Table 5: COVID-19 Negative Reference Product Compliance Test Results
Batch.NO. | 10 measurement results （IgM/IgG）
20170104 | -/- | -/- | -/- | -/- | -/-
| -/- | -/- | -/- | -/- | -/-
20170405 | 10 measurement results （IgM/IgG）
| -/- | -/- | -/- | -/- | -/-
| -/- | -/- | -/- | -/- | -/-
20170502 | 10 measurement results （IgM/IgG）
| -/- | -/- | -/- | -/- | -/-
| -/- | -/- | -/- | -/- | -/-
Three batches of kits meet the requirements for positive / negative reference

## LOD
Table 7：COVID-19 the lowest detection limit (LOD) Test Result（IgM/IgG）
Batch NO.of Reagent | S1 measurement results of three batches |  |  | S2 measurement results of three batches |  |  | S3 measurement results of three batches
20170104 | -/- | -/- | -/- | +/- | +/- | +/- | -/+ | -/+ | -/+
20170405 | -/- | -/- | -/- | +/- | +/- | +/- | -/+ | -/+ | -/+
20170502 | -/- | -/- | -/- | +/- | +/- | +/- | -/+ | -/+ | -/+
According to the results of the above three batches of experiments, the measured values of reference S1 of this kit were negative, and the measured values of S2 and S3 were positive.

## Repeatablility
COVID-19 Repeatability Test Results（IgM/IgG）
Batch NO.of Reagent | S2 measurement results of 10 times |  | S3 measurement results of 10 times
20170104 | +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
20170405 | +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
20170502 | +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
| +/- | +/- | -/+ | -/+
During three batches of repeatability tests of the novel Coronavirus (COVID-19) antibody (IgM / IgG) detection kit (colloidal gold immunochromatography), 10 repeatability tests of S2 and S3 reference products were all positive.

## Specificity
Table 9: Analysis Results of COVID-19 Specificity (IgM/IgG)
Sample | Batch NO. |  |  | Batch NO. |  |  | Batch NO.
| Analysis Results of Specificity |  |  | Analysis Results of Specificity |  |  | Analysis Results of Specificity
Parainfluenza virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Influenza A virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Influenza B virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Chlamydia pneumoniae antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Mycoplasma pneumoniae antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Adenovirus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Respiratory syncytial virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Hepatitis B surface antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Hepatitis C virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Treponema pallidum antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Human immunodeficiency virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
EB virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Measles virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Cytomegalovirus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Enterovirus 71 antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Mumps virus antibody | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Varicella-zoster virus positive samples | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Bilirubin 250 μmol / L | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Hemoglobin 9g / L | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Triglyceride 15mmol / L | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Rheumatoid factor 80IU / mL | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Antinuclear antibody (ANA) 1: 240 | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Anti-mitochondrial antibody (AMA) 80U / mL | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Mouse IgG 1000 μg / mL | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Histamine hydrochloride | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
α-interferon | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Zanamivir | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Ribavirin | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Oseltamivir | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Peramivir | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Lopinavir | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Ritonavir | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Arbidol | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Levofloxacin | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Azithromycin | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Ceftriaxone | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Meropenem | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
Tobramycin | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/- | -/-
According to the above experimental results, there is no cross-reaction and interference to the above substances.

## Interference-detection
Table 10: Interference-detection results of IgM antibody against the disrupted novel coronavirus (covid-19) (IgM/IgG)
Sample NO. | Result | Batch NO. 20170104 | Batch NO. 20170104 | Batch NO. 20170104
|  | Interference detection results | Interference detection results | Interference detection results
Sample 1 | +/- | +/- | +/- | +/-
Sample 2 | +/+ | +/+ | +/+ | +/+
Sample 3 | +/+ | +/+ | +/+ | +/+
Sample 4 | +/- | +/- | +/- | +/-
Sample 5 | +/- | +/- | +/- | +/-
Sample 6 | +/- | +/- | +/- | +/-
Sample 7 | +/+ | +/+ | +/+ | +/+
Sample 8 | +/+ | +/+ | +/+ | +/+
Sample 9 | +/- | +/- | +/- | +/-
Sample 10 | +/+ | +/+ | +/+ | +/+
According to the above experimental results, it can be confirmed the kit is not interfered by the disrupted coronavirus IgM antibody.